## Report
Report
| Child Tag | Text |  | Child Tag | Parent Tag |  | Orphan Tags
0 | PUMP:URS:1 | The pump shall provide both a bolus and basal feature. | 0 | PUMP:URS:1 |  | 0 | PUMP:URS:1
1 | PUMP:URS:3 | The software shall provide a bolus calculator which is supported by programmable correction factors and carb ratios. | 1 | PUMP:URS:3 |  | 1 | PUMP:URS:3
2 | PUMP:URS:8 | The software shall provide a reverse correction feature for the bolus calculator. | 2 | PUMP:URS:8 |  | 2 | PUMP:URS:8
3 | PUMP:URS:10 | The software shall provide a means for the user to select between at least five different European human languages. | 3 | PUMP:URS:10 |  | 3 | PUMP:URS:10
4 | PUMP:URS:100 | The pump shall include a rechargeable Battery. | 4 | PUMP:URS:100 |  | 4 | PUMP:URS:100
5 | PUMP:URS:103 | The pump shall display remaining battery charge. | 5 | PUMP:URS:103 |  | 5 | PUMP:URS:103
6 | PUMP:URS:1000 | The pump shall display remaining insulin volume. | 6 | PUMP:URS:1000 |  | 6 | PUMP:URS:1000
7 | PUMP:URS:3330 | The pump shall be small and lightweight. | 7 | PUMP:URS:3330 |  | 7 | PUMP:URS:3330
8 | PUMP:URS:3350 | The pump shall be easy to use. | 8 | PUMP:URS:3350 |  | 8 | PUMP:URS:3350
9 | PUMP:URS:4000 | The pump shall include an automated dosing algorithm. | 9 | PUMP:URS:4000 |  | 9 | PUMP:URS:4000
10 | PUMP:HRD:100 | Details regarding the rechargeable Lithium Polymer Battery. | 10 | PUMP:HRD:100 | [PUMP:HRS:100]
11 | PUMP:HRD:105 | Details regarding the fuel gauge hardware for the lithium polymer battery. The battery charge shall be displayed to the user. | 11 | PUMP:HRD:105 | [PUMP:HRS:103]
12 | PUMP:HRD:1000 | Details regarding the pressure sensors for use in conjunction with the ideal gas law. | 12 | PUMP:HRD:1000 | [PUMP:HRS:1000]
13 | PUMP:HRD:3330 | Details regarding the size and weight of the pump. | 13 | PUMP:HRD:3330 | [PUMP:HRS:3330]
14 | PUMP:HRD:3350 | Details regarding the full color touchscreen. | 14 | PUMP:HRD:3350 | [PUMP:HRS:3350]
15 | PUMP:HRS:100 | The pump shall include a rechargeable Lithium Polymer Battery. | 15 | PUMP:HRS:100 | [PUMP:PRS:100]
16 | PUMP:HRS:105 | The pump shall include fuel gauge hardware for the lithium polymer battery. The battery charge shall be displayed to the user. | 16 | PUMP:HRS:105 | [PUMP:PRS:103]
17 | PUMP:HRS:1000 | The pump shall include pressure sensors for use in conjunction with the ideal gas law. The gas law shall be used to estimate remaining insulin volume. | 17 | PUMP:HRS:1000 | [PUMP:PRS:1000]
18 | PUMP:HRS:3330 | The pump shall weight no more than 8 ounces dry. | 18 | PUMP:HRS:3330 | [PUMP:PRS:3330]
19 | PUMP:HRS:3340 | The pump shall fit within a volume of 3” by 2” by 0.75”. | 19 | PUMP:HRS:3340 | [PUMP:PRS:3330]
20 | PUMP:HRS:3350 | The pump shall include a full color touchscreen. | 20 | PUMP:HRS:3350 | [PUMP:PRS:3350]
21 | PUMP:HTP:100 | Test 100 | 21 | PUMP:HTP:100 | [PUMP:HRS:100]
22 | PUMP:HTP:200 | Test 200 | 22 | PUMP:HTP:200 | [PUMP:HRS:105]
23 | PUMP:HTP:300 | Test 300 | 23 | PUMP:HTP:300 | [PUMP:HRS:1000]
24 | PUMP:HTP:400 | Test 400 | 24 | PUMP:HTP:400 | [PUMP:HRS:3330]
25 | PUMP:HTP:500 | Test 500 | 25 | PUMP:HTP:500 | [PUMP:HRS:3350]
26 | PUMP:HTP:1100 | Test 1100 | 26 | PUMP:HTP:1100 | [PUMP:HRD:100]
27 | PUMP:HTP:1200 | Test 1200 | 27 | PUMP:HTP:1200 | [PUMP:HRD:105]
28 | PUMP:HTP:1300 | Test 1300 | 28 | PUMP:HTP:1300 | [PUMP:HRD:1000]
29 | PUMP:HTP:1400 | Test 1400 | 29 | PUMP:HTP:1400 | [PUMP:HRD:3330]
30 | PUMP:HTP:1500 | Test 1500 | 30 | PUMP:HTP:1500 | [PUMP:HRD:3350]
31 | PUMP:HTR:100 | Test 100 | 31 | PUMP:HTR:100 | [PUMP:HTP:100]
32 | PUMP:HTR:200 | Test 200 | 32 | PUMP:HTR:200 | [PUMP:HTP:200]
33 | PUMP:HTR:300 | Test 300 | 33 | PUMP:HTR:300 | [PUMP:HTP:300]
34 | PUMP:HTR:400 | Test 400 | 34 | PUMP:HTR:400 | [PUMP:HTP:400]
35 | PUMP:HTR:500 | Test 500 | 35 | PUMP:HTR:500 | [PUMP:HTP:500]
36 | PUMP:HTR:1100 | Test 1100 | 36 | PUMP:HTR:1100 | [PUMP:HTP:1100]
37 | PUMP:HTR:1200 | Test 1200 | 37 | PUMP:HTR:1200 | [PUMP:HTP:1200]
38 | PUMP:HTR:1300 | Test 1300 | 38 | PUMP:HTR:1300 | [PUMP:HTP:1300]
39 | PUMP:HTR:1400 | Test 1400 | 39 | PUMP:HTR:1400 | [PUMP:HTP:1400]
40 | PUMP:HTR:1500 | Test 1500 | 40 | PUMP:HTR:1500 | [PUMP:HTP:1500]
41 | PUMP:PRS:1 | The pump shall provide both a bolus and basal feature. | 41 | PUMP:PRS:1 | [PUMP:URS:1]
42 | PUMP:PRS:2 | The pump shall limit boluses to not exceed 25 units. | 42 | PUMP:PRS:2 | [PUMP:RISK:10]
43 | PUMP:PRS:3 | The pump shall limit boluses to not exceed 15 units/hour. | 43 | PUMP:PRS:3 | [PUMP:RISK:20]
44 | PUMP:PRS:4 | The software shall provide a programmable correction factor feature. | 44 | PUMP:PRS:4 | [PUMP:URS:3]
45 | PUMP:PRS:5 | The software shall provide a programmable carb ratio feature. | 45 | PUMP:PRS:5 | [PUMP:URS:3]
46 | PUMP:PRS:8 | The software shall provide a reverse correction feature for the bolus calculator. | 46 | PUMP:PRS:8 | [PUMP:URS:8] [PUMP:RISK:30]
47 | PUMP:PRS:10 | The software shall provide a means for the user to select between at least five different European human languages. | 47 | PUMP:PRS:10 | [PUMP:URS:10]
48 | PUMP:PRS:100 | The pump shall include a rechargeable Lithium Polymer Battery. | 48 | PUMP:PRS:100 | [PUMP:URS:100]
49 | PUMP:PRS:105 | The pump shall include fuel gauge hardware for the lithium polymer battery. The battery charge shall be displayed to the user. | 49 | PUMP:PRS:105 | [PUMP:URS:103] [PUMP:RISK:40]
50 | PUMP:PRS:1000 | The pump shall include pressure sensors for use in conjunction with the ideal gas law. The gas law shall be used to estimate remaining insulin volume. | 50 | PUMP:PRS:1000 | [PUMP:URS:1000] [PUMP:RISK:50]
51 | PUMP:PRS:3330 | The pump shall weight no more than 8 ounces dry. | 51 | PUMP:PRS:3330 | [PUMP:URS:3330]
52 | PUMP:PRS:3340 | The pump shall fit within a volume of 3” by 2” by 0.75”. | 52 | PUMP:PRS:3340 | [PUMP:URS:3330]
53 | PUMP:PRS:3350 | The pump shall include a full color touchscreen. | 53 | PUMP:PRS:3350 | [PUMP:URS:3350]
54 | PUMP:PRS:4000 | The pump shall include an automated dosing algorithm. | 54 | PUMP:PRS:4000 | [PUMP:URS:4000]
55 | PUMP:RISK:10 | The pump shall deliver no bolus larger than 25 units. | 55 | PUMP:RISK:10
56 | PUMP:RISK:20 | The pump shall deliver no basal rate larger than 15 units/hr. | 56 | PUMP:RISK:20
57 | PUMP:RISK:30 | The bolus calculator shall provide reverse correction. | 57 | PUMP:RISK:30
58 | PUMP:RISK:40 | The pump shall display remaining battery charge. | 58 | PUMP:RISK:40
59 | PUMP:RISK:50 | The pump shall display remaining insulin volume. | 59 | PUMP:RISK:50
60 | PUMP:SDS:10 | Here are details of how the bolus calculator works …. | 60 | PUMP:SDS:10 | [BOLUS:SRS:1] [BOLUS:SRS:2] [BOLUS:SRS:5] [BOLUS:SRS:6] [BOLUS:SRS:8] [BOLUS:SRS:12] [ACE:SRS:1] [ACE:SRS:5] [ACE:SRS:6]
61 | PUMP:SDS:20 | Here are details of how the basal rate works …. | 61 | PUMP:SDS:20 | [ACE:SRS:2]
62 | PUMP:SDS:30 | Here are details of how the PID algorithm works …. | 62 | PUMP:SDS:30 | [AID:SRS:1] [AID:SRS:2] [AID:SRS:10] [AID:SRS:12] [AID:SRS:20]
63 | PUMP:SDS:40 | Here are details of how the GUI works…. | 63 | PUMP:SDS:40 | [ACE:SRS:110] [ACE:SRS:120]
64 | PUMP:SDS:50 | Here are details of how the remaining insulin estimation works… include lots of math …. | 64 | PUMP:SDS:50 | [ACE:SRS:110]
65 | PUMP:SDS:60 | Here are details of how the human languages are swapped around | 65 | PUMP:SDS:60 | [ACE:SRS:10]
66 | PUMP:SDS:70 | Here are details of how the battery estimation works | 66 | PUMP:SDS:70 | [ACE:SRS:100]
67 | ACE:SRS:1 | The software shall provide a bolus feature which generates boluses in the range of 0.01 to 25 units, which an increment of 0.01 units. | 67 | ACE:SRS:1 | [PUMP:PRS:1] [PUMP:TBV:1]
68 | ACE:SRS:2 | The software shall provide a programmable basal rate feature, with basal rates in the range of 0.001 to 15 units/hour in increments of 0.001 u/hr. | 68 | ACE:SRS:2 | [PUMP:PRS:1]
69 | ACE:SRS:5 | The software shall provide a programmable correction factor feature. The software shall support correction factors in the range of 1:10 to 1:600 u/mg/dL. | 69 | ACE:SRS:5 | [PUMP:PRS:5]
70 | ACE:SRS:6 | The software shall provide a programmable carb ratio feature. The software shall support carb ratios in the range of 1:10 to 1:600 u/g. | 70 | ACE:SRS:6 | [PUMP:PRS:6]
71 | ACE:SRS:10 | The software shall provide a means for the user to select between English, German, French, Spanish, and Czech. | 71 | ACE:SRS:10 | [PUMP:PRS:10]
72 | ACE:SRS:100 | The GUI shall indicate percent battery charge. The charge shall be displayed in increments of 1%. | 72 | ACE:SRS:100 | [PUMP:PRS:105]
73 | BOLUS:SRS:1 | The software shall provide a bolus calculator feature. | 73 | BOLUS:SRS:1 | [PUMP:PRS:1]
74 | BOLUS:SRS:2 | The bolus size shall be estimated as the sum of the insulin needed to cover the current meal and correct down to 110 mg/dL. If the sum is negative, the bolus is aborted. | 74 | BOLUS:SRS:2 | [PUMP:PRS:1]
75 | BOLUS:SRS:5 | The amount of insulin needed to cover the meal shall be estimated only when the user indicates a meal was consumed. The volume of insulin shall be estimated as the number of carbs divided by the CarbRatio. | 75 | BOLUS:SRS:5 | [PUMP:PRS:1] [PUMP:PRS:5]
76 | BOLUS:SRS:6 | The amount of insulin needed to correct glucose levels shall be estimated by the following formula if and only if the glucose is greater than 110: | 76 | BOLUS:SRS:6 | [PUMP:PRS:1] [PUMP:PRS:3]
77 | BOLUS:SRS:8 | The value of CorrectionBolus shall be set to zero when the glucose is between 80 and 110 mg/dL, inclusive. | 77 | BOLUS:SRS:8 | [PUMP:PRS:1]
78 | BOLUS:SRS:12 | The amount of insulin needed to correct glucose levels shall be estimated by the following formula if and only if the glucose is less than 80. Please note that in this case, the dose amount is a negative number. | 78 | BOLUS:SRS:12 | [PUMP:PRS:1] [PUMP:PRS:8]
79 | AID:SRS:1 | The dosing algorithm shall include a proportional gain term. The algorithm shall allow proportional gains in the range of 0 to 2.0 in increments of 0.001 | 79 | AID:SRS:1 | [PUMP:PRS:4000] [PUMP:DER:2]
80 | AID:SRS:2 | The dosing algorithm shall include a derivative gain term. The algorithm shall allow derivative gains in the range of 0 to 2.0 in increments of 0.001 | 80 | AID:SRS:2 | [PUMP:PRS:4000] [PUMP:DER:2]
81 | AID:SRS:10 | The dosing algorithm shall include an integral gain term. The algorithm shall allow integral gains in the range of 0 to 2.0 in increments of 0.001 | 81 | AID:SRS:10 | [PUMP:PRS:4000] [PUMP:DER:2]
82 | AID:SRS:12 | The dosing algorithm shall include an integrator window to minimize limit-cycling. The window shall be settable at compile time to a range of 0.002 to 0.01 in increments of 0.001 | 82 | AID:SRS:12 | [PUMP:PRS:4000] [PUMP:DER:2]
83 | AID:SRS:20 | The dosing algorithm shall include anti-windup reset protection. | 83 | AID:SRS:20 | [PUMP:PRS:4000] [PUMP:DER:2]
84 | PUMP:SVAL:100 | This test validates bolus features… blah, blah, blah | 84 | PUMP:SVAL:100 | [ACE:SRS:1] [ACE:SRS:5] [ACE:SRS:6] [BOLUS:SRS:1] [BOLUS:SRS:2] [BOLUS:SRS:5] [BOLUS:SRS:6] [BOLUS:SRS:8]
85 | PUMP:SVAL:200 | This test validates basal features… blah, blah, blah | 85 | PUMP:SVAL:200 | [ACE:SRS:2]
86 | PUMP:SVAL:300 | This test validates the GUI, including human languages… blah, blah, blah | 86 | PUMP:SVAL:300 | [ACE:SRS:10] [ACE:SRS:100] [ACE:SRS:1000] [ACE:SRS:120]
87 | PUMP:SVAL:400 | This test validates remaining volume estimation… blah, blah, blah | 87 | PUMP:SVAL:400 | [BOLUS:SRS:12]
88 | PUMP:SVAL:500 | This test validates remaining volume estimation… blah, blah, blah | 88 | PUMP:SVAL:500 | [AID:SRS:1] [AID:SRS:2] [AID:SRS:10] [AID:SRS:12] [AID:SRS:20]
89 | PUMP:SVATR:100 | Test 100 | 89 | PUMP:SVATR:100 | [PUMP:SVAL:100]
90 | PUMP:SVATR:200 | Test 200 | 90 | PUMP:SVATR:200 | [PUMP:SVAL:200]
91 | PUMP:SVATR:300 | Test 300 | 91 | PUMP:SVATR:300 | [PUMP:SVAL:300]
92 | PUMP:SVATR:400 | Test 400 | 92 | PUMP:SVATR:400 | [PUMP:SVAL:400]
93 | PUMP:SVATR:500 | Test 500 | 93 | PUMP:SVATR:500 | [PUMP:SVAL:500]
94 | PUMP:UT:100 | Unit Test for pump_aaa.c,h | 94 | PUMP:UT:100 | [PUMP:UNIT:100]
95 | PUMP:UT:110 | Unit Test for pump_bbb.c,h | 95 | PUMP:UT:110 | [PUMP:UNIT:110]
96 | PUMP:UT:120 | Unit Test for pump_ccc.c,h | 96 | PUMP:UT:120 | [PUMP:UNIT:120]
97 | PUMP:UT:130 | Unit Test for pump_ddd.c,h | 97 | PUMP:UT:130 | [PUMP:UNIT:130]
98 | PUMP:UT:140 | Unit Test for pump_eee.c,h | 98 | PUMP:UT:140 | [PUMP:UNIT:140]
99 | PUMP:UT:150 | Unit Test for pump_fff.c,h | 99 | PUMP:UT:150 | [PUMP:UNIT:150]
100 | PUMP:UT:160 | Unit Test for pump_ggg.c,h | 100 | PUMP:UT:160 | [PUMP:UNIT:160]
101 | PUMP:UT:170 | Unit Test for pump_hhh.c,h | 101 | PUMP:UT:170 | [PUMP:UNIT:170]
102 | PUMP:UT:180 | Unit Test for pump_iii.c,h | 102 | PUMP:UT:180 | [PUMP:UNIT:180]
103 | PUMP:UT:190 | Unit Test for pump_jjj.c,h | 103 | PUMP:UT:190 | [PUMP:UNIT:190]
104 | PUMP:UT:200 | Unit Test for pump_kkk.c,h | 104 | PUMP:UT:200 | [PUMP:UNIT:200]
105 | PUMP:UT:210 | Unit Test for pump_lll.c,h | 105 | PUMP:UT:210 | [PUMP:UNIT:210]
106 | PUMP:UT:220 | Unit Test for pump_mmm.c,h | 106 | PUMP:UT:220 | [PUMP:UNIT:220]
107 | PUMP:INS:100 | Code Inspection Test for pump_aaa.c,h | 107 | PUMP:INS:100 | [PUMP:UNIT:100]
108 | PUMP:INS:110 | Code Inspection Test for pump_bbb.c,h | 108 | PUMP:INS:110 | [PUMP:UNIT:110]
109 | PUMP:INS:120 | Code Inspection Test for pump_ccc.c,h | 109 | PUMP:INS:120 | [PUMP:UNIT:120]
110 | PUMP:INS:130 | Code Inspection Test for pump_ddd.c,h | 110 | PUMP:INS:130 | [PUMP:UNIT:130]
111 | PUMP:INS:140 | Code Inspection Test for pump_eee.c,h | 111 | PUMP:INS:140 | [PUMP:UNIT:140]
112 | PUMP:INS:150 | Code Inspection Test for pump_fff.c,h | 112 | PUMP:INS:150 | [PUMP:UNIT:150]
113 | PUMP:INS:160 | Code Inspection Test for pump_ggg.c,h | 113 | PUMP:INS:160 | [PUMP:UNIT:160]
114 | PUMP:INS:170 | Code Inspection Test for pump_hhh.c,h | 114 | PUMP:INS:170 | [PUMP:UNIT:170]
115 | PUMP:INS:180 | Code Inspection Test for pump_iii.c,h | 115 | PUMP:INS:180 | [PUMP:UNIT:180]
116 | PUMP:INS:190 | Code Inspection Test for pump_jjj.c,h | 116 | PUMP:INS:190 | [PUMP:UNIT:190]
117 | PUMP:INS:200 | Code Inspection Test for pump_kkk.c,h | 117 | PUMP:INS:200 | [PUMP:UNIT:200]
118 | PUMP:INS:210 | Code Inspection Test for pump_lll.c,h | 118 | PUMP:INS:210 | [PUMP:UNIT:210]
119 | PUMP:INS:220 | Code Inspection Test for pump_mmm.c,h | 119 | PUMP:INS:220 | [PUMP:UNIT:220]